Analysis of Italian P-type detectors

Jason Feldt

In this report, we analyze detectors with positively doped bulk (p-type) and attempt to discover why they much more current than theoretically predicted.  As of yet, we have no culprit, but we have some interesting observations.

figure 1

In figure 1, we see a simple iV curve for detector 182-8.  As this detector breaks down around 80V, it serves nicely as the testbed for the rest of our investigations.  We looked at channels 1, 8, 16, 24, and 32 to see if there was a distribution of individual strip current as a function of position along the detector.  There was no appreciable deviation from flatness.  That is, individual strips had currents roughly the same current, an expected 1/32 of the backplane current. Thus there were no "hot" (high current carrying) channels.  	We picked strip 8 and plotted the ratio of backplane current to its current in figure 2.  Other than in the breakdown region, strips 8 and 9 have a ratio is in the 20's, which is close to the expected 32, and strip 7 is even closer to 32.

figure 2

We decided to hone our efforts on strips 7,8, and 9.  We checked the capacitors on these channels to make sure worked properly, meaning the current should remain small and not change appreciably while sweeping the AC pad voltage relative to the strip voltage.  The results showed that the capacitors were intact, but not without some wrangling on our part.  The data shown in figure 3 looks suspect and intially caused great excitement, but we ran the sweep of the AC pad voltage in reverse and discovered that the variation in strip voltage was due to charging currents.  In figure 4, one can see that the forward and backward sweep currents are mirrored across 0V.

figure 3

figure 4

figure 5

Figure 5 shows iV curves on a logarithmic scale for backplane sweeps while holding the AC pads at various voltages.  Other than breakdown behavior, little changes between AC pad biases of –1V, 0V, and 1V.

In the meantime, results such as those shown in figure 6 confirmed that the effect is magnified when more strips are subjected to an AC pad bias.  We decided to tie together all the AC pads in order to magnify any effect that may have been hidden in the noise previously.  Before the bonding, we checked the capacitor on each channel for excess current.  The results in figure 7 checked out, so we bonded the AC pads together.

figure 6

figure 7

The measurements taken after bonding show an increase in effect in the iV curve in figure 8, where the breakdown behavior is cut off in order to show the relevant data more clearly. At 70V, in order of largest currents to smallest are –2V, -1V, 2V, 1V, 0V.  This is not as correlated as it seems it should be.  And, there is a crossover around 23V.  This mysteriousness could be explained by charging currents (?), which are transient and die off after an initial startup period.	The effects of these charging currents are again demonstrated in figure 9.  Notice that the variation in current with respect to AC pad voltage is less than in the three strip case in figure 4.  This seems strange as well and is a source for further investigation.

figure 8

figure 9

Data was taken over a period to let any transient behavior settle down.  The transient behavior is mostly clearly seen at lower backplane voltages, which we expect, as we begin with low voltages as we ramp up our iV graphs.  Figures 10,11, and 12 show timebase data for AC pad voltages of –2V, 0V, and 2V.

figures 10 & 11

figure 12