How to Test a GLAST Ladder

June 20, 1999

M.Hirayama

(J:\glast\tracker\doc\HowToTestLadder.doc)

# Preparation

## Setup Instruments

Connect all cables from the backside prober to Keithley 238 through a coax-to-triax conversion box.

For cap-test, also connect Hpot to Hcur, Lpot to Lcur of the LCR meter HP4284.

## Invoke icv_curve.vi

Double click icv_curve.vi in “Programs for Testing” directory on the desktop.

Actual file is J:\glastlab\labview\detectorTest\icv_curve.vi.

Set values for “Voltage Scan” as follows:

|  | For Detector | For Ladder |
| --- | --- | --- |
| GPIB address | 12 | 12 |
| Start voltage | 10.0 V | 10.0 V |
| Stop voltage | 200.0 V | 200.0 V |
| Voltage step | 10.0 V | 10.0 V |
| compliance | 1 A | 20 A |
| Measurement waiting time | 1 s | 5 s |

Set “Constant Voltage Supply” to “Don’t Use”.

Set “LCR meter” to “Don’t Use”.

## Invoke c_auto2.vi (for cap-test only)

Double click c_auto2.vi in “Programs for Testing” directory on the desktop.

Actual file is J:\glastlab\labview\detectorTest\c_auto2.vi.

Set values for “LCR meter” as follows:

| GPIB address | 3 |
| --- | --- |
| Function Type | CSRS (6) |
| Frequency | 1.0E+3 (1 kHz) |
| Voltage Level | 1 V |

Also, set values on the panel as follows:

| How many strips | 320 |
| --- | --- |
| X-step | 0.0 |
| Y-step | -0.194 |

## Open Databases

Open ladders.fp3 with FileMaker Pro and select a layer shown in a table below.

For post-glue IV measurement, open detectors.fp3 with FileMaker Pro and select a layer shown in a table below.

|  | Layout in ladder database “ladders.fp3” to open | Layout in detector database “detectors.fp3” to open |
| --- | --- | --- |
| Post-glue IV | Ladder | PostGlue Current Entry |
| Post-bond IV & cap-test | CapTest Entry | None |
| Post-pot IV | PostPot Current Entry | None |

# Test Proceedure

## Load a Ladder onto Probe Station

Load a ladder to be tested on the chuck of the probe station as follows:

- Put a vacuum tube on the fixture.  Turn on the vacuum.

- Put a fixture on a backside prober.  Use alignment pins to guide it.

- Put the fixture on a dove-tail holder on a chuck of the probe station together with the backside prober.  Be sure not to drop the backside prober when you lift the fixture.

- Tighten screws on the sides of the dove-tail holder by hand.

- Remove the cover of the fixture.

## Put a Probe Down on a Bias Pad

Locate a bias pad of the detector and probe it with a steel probe as follows:

- Make sure there is enough clearance under a stage of the probe station and under a probe tip and press “Home” on LabMaster.

- Locate a bias pad depending on your measurement as follows:

| Measurement | Bias pad used |
| --- | --- |
| Post-glue IV | A bias pad on the bias resistor side of detector A in a ladder, next to strip 1 |
| Post-bond IV & Cap-test | A bias pad on the electronics side of detector A in a ladder, next to strip 1 |
| Post-pot IV | Same as above |

- Bring the bias pad roughly at the center of a monitor.

- Press “Fine Up” on LabMaster.

- Put a probe tip down on the bias pad.  Press “Find Down” then “Fine Up”, and confirm a good contact through a microscope on the probe station.

- Press “Fine Down” and put a cover on the Al box.

- Press “Fine Up”.  Check a good contact through the monitor.

## Do Your Measurement

Actual procedure of measurement depends on a production step.  Perform 2.3.1, 2.3.2, or 2.3.3 depending on the production step you are in.

### Post-glue Detector IV Measurement

- Turn off all the lights in the box.

- Find the ladder being tested in the ladder database and click one of detectors in a picture of a ladder on the top-right corner.  The detector database comes up.  Press “Sort” on the page.

- On a window of icv_curve.vi, choose “Detector” to measure, put detector ID, temperature, and humidity.

- Run icv_curve.vi.  You will be asked to select a directory and a filename to put measurement results are stored.  Use a filename as suggested by the program.  Choose a directory “J:\glast\tracker\PrototypeTower1\detector\Ln”, where n is a ladder ID of the ladder to be measured, and press “Save” to start a measurement.

- Write down currents at 100 V and 150 V in a log book.  Put the numbers into the detector database. Put date and your initial on the same line. Choose “accepted” or “rejected” depending on the measurement result.

- Press “Fine Down” and move the chuck to the next detector.

- Adjust XY so that a probe tip can touch a bias pad of the detector.

- Press “Fine Up”.  Check a good contact through the monitor.

- Repeat 14) to 19) until all detectors in a ladder are measured.

### Post-Bond Ladder IV Measurement and Capacitance Test

- Turn off all the lights in the box.

- On a window of icv_curve.vi, choose “Ladder” to measure, put ladder ID, temperature, and humidity.

- Run icv_curve.vi.  You will be asked to select a directory and a filename to put measurement results are stored.  Use a filename as suggested by the program.  Choose a directory “J:\glast\tracker\PrototypeTower1\detector\Ln”, where n is a ladder ID of the ladder to be measured, and press “Save” to start a measurement.

- Write down currents at 100 V and 150 V in a log book.  Put the numbers into the ladder database. Put date and your initial on the same line. Choose „accepted“ or „rejected“ depending on the measurement result.

- Press “Fine Down”.

- Change cable connections for capacitance test and turn on all the lights in the box.

- Align strips of detector A to X-axis of chuck motion.

- Move the chuck to probe strip 1.

- Press “Fine Up”.  Check a good contact through the monitor.

- On a window of c_auto2.vi, put ladder ID, temperature, and humidity.

- Run c_auto2.vi. You will be asked to select a directory and a filename to put measurement results are stored.  Use a filename as suggested by the program.  Choose a directory “J:\glast\tracker\PrototypeTower1\detector\Ln”, where n is a ladder ID of the ladder to be measured, and press “Save” to start a measurement.

- Capacitance you have just measured should be about 1.8 nF for all strips.  If you find a coupling capacitance of other than 1.8 nF, write it down on a log book and put a comment on a ladder data sheet.  Put the comment into the ladder database.  Locate a strip with a problem and repair it if necessary.  See a table below for typical cases.

- Resume cable connections for current measurement.

| Falure mode | Coupling capacitance | What to be done |
| --- | --- | --- |
| Coupling shourt | A couple of F | Remove wire(s) to the shorted detector |
| Al strip break | Less than 1.8 nF | Locate the break, take a picture of the break, and store it in a directory “photo” |
| Al strips shorted | 3.6 nF, 5.4 nF, or multiplet of 1.8 nF | Locate the short and repair it. If impossible to repair, remove the closest wires on the shorted strips. |
| Missing wire | 1.2 nF or 0.6 nF | Locate the wire and repair it. Measure coupling capacitance of the strip after the repair. |

### Post-Pot Ladder IV Measurement

- Turn off all the lights in the box.

- On a window of icv_curve.vi, choose “Ladder” to measure, put ladder ID, temperature, and humidity.

- Run icv_curve.vi.  You will be asked to select a directory and a filename to put measurement results are stored.  Use a filename as suggested by the program.  Choose a directory “J:\glast\tracker\PrototypeTower1\detector\Ln”, where n is a ladder ID of the ladder to be measured, and press “Save” to start a measurement.

- Write donw currents at 100 V and 150 V in a log book.  Put the numbers into the ladder database. Put date and your initial on the same line. Choose “accepted” or “rejected” depending on the measurement result.

## Unload the Ladder

Unload the ladder from the probe station as follows:

- Press “Gross Down”.

- Double click anywhere on a video image to bring a window titled “AWP1080ATC”.  Press “Unload”.

- Remove a cover from the Al box.

- Put a cover on the fixture and screws it down.

- Loose the screws on the side of the dove-tail holder and remove the fixture and the backside prober from the dove-tail holder.

- Carefully lift the fixture on the backside prober to separate them.

- Turn off the vacuum and remove a vacuum tube from the fixture.

## Update the Ladder Database after Post-Glue measurement

If you are measuring post-glue currents of detectors in a ladder, update the ladder database as follows. Otherwise, the database has already been updated during the measurement procedure described above.

- On the detector database, press “Go to Ladder Database”.  The ladder database comes up.

- Read summed currents at 100 V and 150 V from the page shown up.  They are located right below a word “Survey” on the page.  Write down the numbers on a paper version of the data sheet.

- Put date and your initial on the left.

- Click “accepted” or “rejected” depending on measurement result.

- Store the fixture with the data sheet paper for wire-bonding.

# After Measurements

After all measurements in a day, you should keep testing area clean and secure. Check following points before you leave the room.

- Put a protection on the back side prober. Use empty fixture to cover it.

- Turn all the lights off in the Al box. Put a cover on the Al box.

- Turn off instruments.

- Minimize a LabMaster window. Do not exit from LabMaster.

- If you are the last person to leave the room, turn the room lights off and lock the door.