Initial Electrical Test of Tray

(J:\glast\documents\Initial Tray E-Test Procedure.doc)

Masaharu Hirayama

#### SCIPP

9/3/99

In each step of this test, a certain result is expected. The expected results are underlined in the description below. If you find a result other than expected, stop the test immediately and report it to one of the GLAST staffs.

### Setup

  - Turn on an HV source Keithley 2410 (or equivalent instrument).

  - Set the HV source as voltage source with current measurement.

  - Connect HV outputs of Keithley to probes: positive side to a steel probe with micro-manipulator, negative side with a handy probe for DVM.

  - Set output voltage to 1 V and current compliance to about 100 uA.

  - Output the voltage by pressing “ON” (or “Operate”, depending on the instrument you  are using).

- Check a dead short to closeout

  - Touch the tray closeout with the steel probe.

  - Touch the tray closeout with the handy probe. The current must reach its compliance.

  - Touch bias pads on the Kapton sheet.  The current must stay less than 10 nA.  You should check at least 5 pads, one pad for each detector ladder.

  - Touch the GND strip on the Kapton sheet.  The current must stay less than 10 nA.

##### Check a dead short to GND plane of Kapton sheet

  - Touch the GND strip of the Kapton sheet with the steel probe.

  - Touch the GND strip of the Kapton sheet with the handy probe. The current must reach its compliance.

  - Touch bias pads on the Kapton sheet.  The current must stay less than 10 nA.  You should check at least 5 pads, one pad for each detector ladder.

  - Touch the tray closeout.  The current must stay less than 10 nA.

##### Check HV Rating of Kapton Sheet

  - Set the HV source voltage to 100 V.

  - Touch bias pads on the Kapton sheet.  The current must stay less than 50 nA.  You should check at least 5 pads, one pad for each detector ladder.

  - Set the HV source voltage to 1 V.

##### Test the other side of the tray

  - Repeat section 2 through 4 on the other side.

  - Update the tray database.