SPECIFICATIONS for GLAST TRACKER

AUTOMATIC WEDGE BONDER

# October 1, 1999

Gwelen Paliaga

### Santa Cruz Institute for Particle Physics

Eduardo do Couto e Silva

### Stanford Linear Accelerator Center

This is a list of the general requirements and preferred features of the wire bonders that will be used to build the silicon tracker for GLAST.

GLAST Components:

- HDI  -  40 cm. X 2.5 cm. High density PC board/Kapton hybrid carrying 30 bare die

3600 bonds

- Detector Ladders  -  4 pieces: 40 cm. X 10 cm. Detector assemblies that need to be wire bonded together.

1350 bonds

- Trays  -  40 cm. X 40 cm. Assembly bonded along one edge.

1800 bonds

## GENERAL

- 100 micron pitch

- 1-2 mil wire

- high frequency ultrasonic generator (100-120 kHz)

- 45, or 60 degree clamp mechanisms  (needs clarification from detailed look at HDI and corner post obstructions)

- Use of standard consumables (wedge and wire) (need to define standard)

- Piezoelectric clamp

- > 360 degree bond head rotation

- GZ or MZ Leica microscope (or equivalent) with video port.  (10 –100x zoom)

## STAGE & MECHANICAL

- Minimum access (throat) 5 in.

- Preferred access 8 in.  Reaches center of the tray for possible ladder repair.

- Minimum bonding area 4 in. x 4 in. (requires software interface to an automatically indexing stage)

- Preferred bonding area 8 in. x 16in.  (larger bonding area reduces #of indexing movements:  16 in. “x” motion would need no indexing)

- Needs to have sufficient platform area to mount indexing work holders for              the 16 in. X 16 in. tray

- Linear encoders required

- Stage must support a minimum of 30 lbs. (super GLAST tray + workholder).

SOFTWARE

- Programmable up to 4000 bonds and 32 systems.

- CAD programming desired (or equivalent offline programming capability)

- Ability to add or remove selected wires from each program run

- Graphical programming  (user friendly for many operators and programmers)

- I/O capabilities such as :  -    error reporting

- diagnostic reporting for analysis (parameters, squash, missing wires, missed bonds, re-bonds)

- interface to peripheral instruments (indexing work holders, electronics testing)

- Missing wire detector required

- Pattern recognition system required

CALIBRATION/ TARGETING

- Repeatability 1-2 micron (Needs research)

- Tool calibration to 3 micron for all orientations (needs research)

SERVICE & SUPPORT

- Maintenance contract

- Application support

- Service and repair turnaround time

- Training