# Plateauing Procedure for the

# QNET DAQ series 6000

By Steve Kliewer

## Plateauing the counters

- Connect all parts together.

  - Stack all 4 detectors and orient horizontally

  - Set all tube voltages on 4-way power supply (PS) to minimum (~0.3 V, fully CCW)

  - Connect all cables before applying power.  Double check.  Then apply power to DAQ.

  - Initialize communications to the DAQ via Hyperterminal.

- Set starting values:

  - Type HE for help and V1 to display current settings.

  - Set the discriminator voltage (VD) for all 4 counters on the QNET board to 0.300 V  (Tl 4 300).  This can be measured between TPGND1 and TPVTH 0, 1, 2, & 3 on the DAQ board.

  - Set parameters on the QNET board for singles counts on all counters, with

    - RE, CD, WT 01 00, WT 02 02, WC 00 0F

    - Check all settings using V1

  - Vary the tube voltage (VT) to determine the range of voltage to produce zero counts up to no more than 200 Hz (~ 0.8 – 1.2V).  VT is measured at the tube power supply (PS).  Use the on-board display scaler and the reset button next to it to measure this rate.  Keep all tube voltages except the one being tested, adjusted to the minimum setting.

  - Set one tube of each pair (Reference counters) so that it yields about 40 Hz of singles counts.

  - Set parameters on the QNET board for 2-fold coincidence counts, d = 40 ns, and w = 100 ns.

- Pair up the counters as Reference and Test as follows

  - Step 1      Ref: Ch 0     Test: Ch 1

  - Step 2      Ref: Ch 1     Test: Ch 0

  - Step 3      Ref: Ch 0     Test: Ch 2

  - Step 4      Ref: Ch 0     Test: Ch 3

- Open the Excel spreadsheet: Plateau Procedure.xls.  Use the command V1 to display and record the following parameters and settings:

  - Date, QNET DAQ serial #, coincidence level, d, w, layout, & timing interval

- Plan the sequence of measurements for each channel

  - Determine a test time interval that will give you a large enough count (N) that the uncertainty ((N) will be small enough (~10%) but not so large that the measurements take too long.  This will probably be 1 or 2 minutes.

  - Vary only the tube voltage designated as Test.  Change the voltage carefully and slowly.  Wait until the count rate stabilizes after every change.

- Measure the counts at several tube voltages for the Test counter.

  - Use the command sequence:

    - CD, RB (start clock), wait for end of test interval, DS, 
CD stops the scrolling of event lines, RB reset the scaler registers, and DS displays the scalers.  Scalers 0, 1, 2, & 3 contain the singles rates for each of the channels, Scaler 4 contains the coincidence count.

    - Be sure to accurately hit enter on DS exactly the test interval after hitting enter on RB.

  - Record this data into the spread sheet

    - Time of day at start.

    - Tube voltage of Test counter.

    - Scaler counts.  These registers read in Hexadecimal.  Record the readings for both counters as well as S4 (coincidence count) in hex into the spreadsheet.

    - The spreadsheet will immediately calculate the singles count rates, the coincidence rate, the uncertainties in each, as well as the probable accidental coincidences as a percent of actual coincidence rate.

    - Observe the Charts of these values after each measurement.

  - Locate the plateau region for the tube under test.

- Repeat paragraph 6 for each Step of the test (see paragraph 3).

- Save the spreadsheet as documentation of this process.

- Acquire data overnight with all counters collecting data at optimum (plateaued) settings in 2-fold coincidence mode.  Upload to eLab and run Performance test.  Look for equal performance between channels with only one peak in frequency vs Time over Threshold graph.

# Text boxes

operating point

plateau

knee